## Slide 1
In-depth Spectrum analysis
Andrea Corrado
August 2024

## Slide 2: Outline
Correction of previous declarations about data
Why delve into spectral analysis
  Expectations on EG spectrum
Processing pipeline to estimate spectrum
  Filtering strategy
  Spectrum estimation strategy
Results on external data
Conclusions

## Slide 3: Outline
Correction of previous declarations about data
Why delve into spectral analysis
  Expectations on EG spectrum
Processing pipeline to estimate spectrum
  Filtering strategy
  Spectrum estimation strategy
Results on external data
Conclusions

## Slide 4: Spare1 trace ≠ Reference trace
So far, we trusted that the reference trace and the spare1 trace were composed by the same signals. But:
Reference
Spare1
Into each MAP, reference and spare1 trace aren’t the same, except for subjects 2 and 4.
Such observation support the option that spare traces have a different nature from reference ones (i.e., unipolar derivations?)
Knowing the true nature of these signals now becomes crucial.
Mean ± SD plots (same Map, different subjects)

## Slide 5: Spare1 trace ≠ Reference trace
Moreover, reference trace signals, in some cases, presents some strange behaviors:
Whole P-QRS-T cycle squeezed in a window shorter than 0.2 seconds
  Such behavior suggest am heart rate higher than 60 bpm, but during the 1-second acquisition there is only a single beat
  These traces present a sort of unclear oscillation temporally located where spare traces suggest the presence of a QRS.
There is often a single beat and in general centered on 0.5 seconds
It’s fundamental knowing the whole process of Reference acquisition to better contextualize these behaviors.

## Slide 6: Outline
Correction of previous declarations about data
Why delve into spectral analysis
  Expectations on EG spectrum
Processing pipeline to estimate spectrum
  Filtering strategy
  Spectrum estimation strategy
Results on external data
Conclusions

## Slide 7: Why delve into spectral analysis
Previously, we found reference traces with unclear behaviours:
  Absence of low-frequency components
  Huge presence of high frequency components
Or:
  Frequency components totally squeezed into low frequences, without any peak
Are these characteristics acceptable? Which is the expected shape of a spectrum of an ECG signal? [1]
  T wave localised before 10H, Highest peak
  P wave localized before 15 Hz, possibly masked by other components
  QRS between 1 and 40 Hz, with one or more peaks.
  NB: reference figure can be used even with AR spectrum estimation. In this case, a single beat manually modified.
Spectrums evaluated so far are not acceptable
  NB: Reference figure obtained using Welch Spectrum Estimation
Sörnmo, Leif, and Pablo Laguna. "Electrocardiogram (ECG) signal processing." (2006).

## Slide 8: Outline
Correction of previous declarations about data
Why delve into spectral analysis
  Expectations on EG spectrum
Processing pipeline to estimate spectrum
  Filtering strategy
  Spectrum estimation strategy
Results on external data
Conclusions

## Slide 9: Processing pipeline
So far, data are taken as they are, without any type of preprocessing
Is known that ECG requires preprocessing because affected from various sources of noise [2]
  Physiological, Environmental, Artifacts and Electronic
There are many strategies used to clean ECG signals: [2]
  Numerical Filters
  Averaging
  Wavelet filtering
  …
2. Saraiva, João & Plácido da Silva, Hugo & Fred, Ana. (2022). Denoising and Artifact Removal of the Electrocardiogram, Electrodermal Activity and Accelerometery for Continuous Ambulatory Monitoring of Epileptic Seizures with Wearable Devices.

## Slide 10: Filtering strategy
Strategy 1: Zero phase band-pass filter [2]
Two Butterworth Zero-Phase filters are used to clean the signal.
Strategy 2: Wavelet thresholding and band-pass filter [2]
Wavelet Soft thresholding to reduce the effect of the high and low frequency noise, improving BP filtering [2,3,4]
BP (zero-phase) filter to remove the residual high frequency noise and remove the possible low-frequency drift [2]
Original signal
DWT decomposition
Donoho soft thresholding
Symmetric padding
De-padding and DWT reconstruction
Butterworth HP filter (order 6, cutoff at 0.5 Hz)
Butterworth LP filter (order 6, cutoff at 45 Hz)
Improves the overall wavelet transformation but involves a phase shift into the de-padded signal[3]
Uses noise variance estimation to fix a threshold under which DWT coefficients are set to zero: where M is the sample size [4]
Removes low frequency drift [2]
Removes High frequency noise residuals [2]
Symlets 4 wavelet is used due to its good performances on ECG signals [2, 3]
3. Singh, Brij N., and Arvind K. Tiwari. "Optimal selection of wavelet basis function applied to ECG signal denoising." Digital signal processing 16.3 (2006): 275-287.
4. Donoho, David L. "De-noising by soft-thresholding." IEEE transactions on information theory 41.3 (1995): 613-627.

## Slide 11: Filtering strategy: wavelet filters
The DWT decomposes a signal into multiple scales, also called levels. At each level it captures both the approximate and detailed components of the signal [3] :
Approximation (LPF) coefficients represent the signal at a coarser resolution (low-frequency content).
Detail coefficients (HPF) represent the finer details (high-frequency content).
The primary idea is that noise predominantly affects the high-frequency components. By thresholding the detail coefficients, one can effectively reduce the noise while preserving the signal.
Soft thresholding [4] (introduced by Donoho) not only sets small coefficients to zero but also shrinks larger coefficients toward zero by the threshold value, smoothing the signal.

## Slide 12: Filtering strategy: padding
To improve the DWT, specially when the number of points of the signal is low (i.e., single beat) ,padding is often used [2]:
Zero padding
Symmetric padding
Circular padding
…
Among all possibilities, it seems that symmetric padding performs better [2,3]. So, in the proposed filtering strategy there is the possibility of using such type of padding.
When padding is present:
  Single-beat filtering performance improved
  The denoised signal is shifted respect to the original
When padding is absent:
  Phase shift is absent
  Presence of artifacts into the denoised signal
It should be noted that, when the number of points increases, the padding is not required. So, from now on, it will not be used.
Padding ON, 1000 points, HF noise
Padding OFF, 1000 points, HF noise
Padding OFF, 2000 points, HF noise

## Slide 13: Filtering strategy 1 or 2?
The higher is the level of cleanliness of the signals, the higher will be the goodness of the spectrum, so strategy 1 (BP zero phase filters) and strategy 2 (wavelet and BP zero phase filters) are compared in terms of performances on two “didactics” signals corrupted by high and low frequency noise.
High frequency noise
Strategy 2 performs better than strategy 1, even if strategy 1 reaches good results
Low frequency noise
Both strategies seems to perform well, with very low differences between them.
Because of:
Goodness of results on filtering
Known higher performance even on other sources of noise [2]
Good performances on short signals [2,3]
Strategy 2 (Wavelet thresholding and BP zero phase filters) will be used from now on.

## Slide 14: Spectrum estimation
Spectrum estimation is made through AR model [5]
Spectrum of the target signal
Spectrum of ARMA model
Spectrum of the input signal of the model (white noise)
Transfer function of the AR model which approximates the ARMA one
ARMA models can be approximated by an AR with sufficiently high order
Fast and easy application
Require a zero-mean signal as input of the estimator of the model
Consistent and not windowed estimation of the spectrum
Order choice is crucial

## Slide 15: Spectrum estimation: AR identification
When dealing with AR models, and in particular if used to evaluate spectrums, the estimation method must be chosen carefully. From previous studies [5,6] is known that:
Yule-Walker estimation:
  Ensure stable models
  Requires high order models and does not have good frequency resolution
Least-Squares estimation:
  Does not ensure stable models (but it’s not a problem if the order is low)
  Have high frequency resolution and performs well with lower orders
Burg’s estimation:
  Ensure stable models
  Ensure high frequency resolution
NB: the difference is mainly on the value (and number) of the coefficients, while the spectrum estimation formula remains the same.
Yule- Walker
Where:
is the vector of AR coefficients (order p)
is the autocovariance matrix (pxp)
is the vector of the first p points of the autocorrelation function
Least squares
Where:
is the vector of AR coefficients (order p)
is matrix of observations lagged by p (N-p x p )
is the vector of the N-p observations
Burg iterative process minimizes
Where:
is the forward prediction error (how good the model predicts future points knowing the past ones)
is the backward prediction error (how good the model predicts the past points knowing the future ones)
Yule- Walker
Where:
is the vector of AR coefficients (order p)
is the autocovariance matrix (pxp)
is the vector of the first p points of the autocorrelation function
Least squares
Where:
is the vector of AR coefficients (order p)
is matrix of observations lagged by p (N-p x p )
is the vector of the N-p observations
Burg iterative process minimizes
Where:
is the forward prediction error (how good the model predicts future points knowing the past ones)
is the backward prediction error (how good the model predicts the past points knowing the future ones)
5. Übeyli, Elif Derya, Dean Cvetkovic, and Irena Cosic. "AR spectral analysis technique for human PPG, ECG and EEG signals." Journal of Medical Systems 32 (2008): 201-206.
6. Nayak, Jagadish, et al. "AR modeling of heart rate signals." 2004 IEEE Region 10 Conference TENCON 2004.. IEEE, 2004.

## Slide 16: Spectrum estimation: AR order choice
AR spectrum estimation is a well-established strategy but requires to fix the order of the model [5]
There are many strategies [6]:
  AIC criterion
  BIC criterion
  AIC/BIC mixed strategy
  Similarity between AR spectrum estimation and classical/Welch spectrogram
But there isn’t a “fixed” rule. In other studies, the optimal order is found to be between 8 and 20 (Burg and LS estimations) [5,6] and all cases was fixed.
Normalized AIC criterium has been used on a range of candidate orders, with a threshold to decide whether the order reached improves significantly the estimation.
p_opt = evaluate_order(signal, min_order, max_order, step, eps, type)
Initialize AIC_vector to store AIC values for each model order
Initialize AIC_difference as an empty array
Initialize p_opt (optimal order) to 0
Loop through each model order in the range of p (min_order:step:max_order):
Fit AR model to the signal (after removing the mean), using the type estimation, for the current order p(i)
Calculate the AIC for the current AR model and store it in AIC_vector[i]
If this is not the first iteration (i > 1):
  Compute the absolute difference between the current AIC and the previous AIC and store it in AIC_difference
  If the AIC difference is less than a tolerance value (eps):
    Break the loop (convergence achieved)
Find the position of the optimal order where AIC_difference is less than epsilon
If no convergence was found (position not found):
Set the optimal position to the index of the minimum AIC value in AIC_vector
Set p_opt (optimal order) to the value of p at the optimal position

## Slide 17: Spectrum estimation: orders boundaries
It’s important decide the range into which found the optimal order.
From literature, an order between 8 and 20 should be enough for LS or Burg estimators. [5,6]
Order choice is computationally demanding
LS estimation seems to be more robust on high orders
Both tent to be more affected by noise when the order increases (>18)
Optimal order was searched between 8 and 16 at the end
Burg estimation
LS estimation

## Slide 18: Outline
Correction of previous declarations about data
Why delve into spectral analysis
  Expectations on EG spectrum
Processing pipeline to estimate spectrum
  Filtering strategy
  Spectrum estimation strategy
Results on external data
Conclusions

## Slide 19: Spectrum estimation: AR order choice
Before applying the algorithm to the AVNRT data, it has been tested on external data:
  Didactical ECG with High Frequency noise (Fc: 1000 Hz, ~ 8500 points)
  Didactical ECG with Low Frequency noise (Fc: 1000 Hz, ~ 9000 points)
  PhysioNet Database ECG data: healthy and pathological (Fc: 360 Hz, 9000 points)
For all these scenarios have been plotted the AR spectrums compared with the Welch spectrum and the FFT periodogram
The number of points onto evaluating the spectrums are increased to see the behaviour of the analysis when considering an increasing number of cardiac beats
Spectrogram
Welch Spectrogram
Where:
is a window of size M
P is the average of the window , defined as
Windows divide the signals in S, partially overlapping, segments

## Slide 20: ECG with High Frequency Noise
Spectrum evaluation performance
Good frequency resolution
Appreciable peaks when expected
Not considerable differences while augmenting the number of points
Convergence of LS and Burg estimations.

## Slide 21: ECG with Low Frequency Noise
Spectrum evaluation performance
Good frequency resolution
Appreciable peaks when expected
Differences in peaks presences and amplitudes
Again, convergence between methods

## Slide 22: PhysioNet database
Healthy record
Good estimation of Welch spectrum even on “real world” data
Peaks presence limited between 0 and 30 Hz
Pathological Record
Good estimation of Welch spectrum
Presence of peaks above 30 Hz, an aspect that can be traced back to the pathology

## Slide 23: Outline
Correction of previous declarations about data
Why delve into spectral analysis
  Expectations on EG spectrum
Processing pipeline to estimate spectrum
  Filtering strategy
  Spectrum estimation strategy
Results on external data
Conclusions

## Slide 24: Conclusions
This presentation shows a method for estimating the spectrum of an ECG signal. In particular:
A mixed filtering strategy, with wavelet thresholding and numerical filter, is used with good performances. Filtering pipeline could include padding, which improves performances on short-length signals but leads to a phase shift.
AR optimal order was found between 8 and 16, while using nAIC as metric.
Least Squares and Burg’s estimation approaches are both used for their high frequency resolution. When augmenting the number of points, provides the same spectrum estimation.
Overall performances of the method are satisfying. In particular, the differences between the spectra of healthy and pathological ECG signals can be seen

## Slide 25: References
Sörnmo, Leif, and Pablo Laguna. "Electrocardiogram (ECG) signal processing." Wiley encyclopaedia of biomedical engineering (2006).
Saraiva, João & Plácido da Silva, Hugo & Fred, Ana. (2022). Denoising and Artifact Removal of the Electrocardiogram, Electrodermal Activity and Accelerometery for Continuous Ambulatory Monitoring of Epileptic Seizures with Wearable Devices. 10.13140/RG.2.2.10053.12004.
Singh, Brij N., and Arvind K. Tiwari. "Optimal selection of wavelet basis function applied to ECG signal denoising." Digital signal processing 16.3 (2006): 275-287.
Donoho, David L. "De-noising by soft-thresholding." IEEE transactions on information theory 41.3 (1995): 613-627.
Übeyli, Elif Derya, Dean Cvetkovic, and Irena Cosic. "AR spectral analysis technique for human PPG, ECG and EEG signals." Journal of Medical Systems 32 (2008): 201-206.
Nayak, Jagadish, et al. "AR modeling of heart rate signals." 2004 IEEE Region 10 Conference TENCON 2004.. IEEE, 2004.